Ex Vivo Scan (µCT 35/vivaCT 40)

Prepared by:  	Helen Buie

Date:		2012-JUN-21

Before the scan day:

| - Register samples for scanning. - see the sample naming convention - register samples using the µCT-Samples program - record sample names and numbers on your log sheet |  |
| --- | --- |
| - Setup a control file. Use the planning document to estimate your scan time and file size. |  |
| - Book the µCT 35 or the vivaCT 40 scanner. |  |

On the scan day:

| - If necessary, turn on the scanner (µCT 35/vivaCT 40) and open the scanning program to start the 20 min warm up. |  |
| --- | --- |
| - Prepare the sample(s) in the ex vivo holder and place the holder in the scanner. - try to orient your samples consistently in the holder - if multiple samples are placed in the tube, make notes of the order of the samples in the holder |  |
| - Perform a scout view. - if it is the first scan - login to the workstation - put in your operator number or name - start the scanning program - select the control file - select the sample number - perform a scout view |  |
| - Place a reference line and submit scan (interactive/batch). - use the batch mode if you have multiple samples in the holder and for long scans (avoids the possibility of someone accidentally cancelling the scan) |  |
| - Record the measurement number on the log sheet. |  |
| - Repeat steps: - for batch scans (multiple samples in the holder or multiple scan sites of a single sample), select a new sample number and repeat steps 7-8 - for new samples, repeat steps 5-8 |  |

| - Post-scanning procedures. - disinfect the ex vivo sample holder or wash with dish soap, as appropriate - close the scanning program - if performing batch scans, request an email notification when the scans are done - in the terminal $ @com:email - log out of the workstation |  |
| --- | --- |